Michigan Department of Health and Human Services

Bureau of Laboratories

AZ.036 Lead in Dust Wipes, Soil and Paint

Rev 03/18/2026

ANALYTES TESTED:

Lead.

USE OF TEST:

Analysis of dust wipe, paint or soil for presence of lead.

SPECIMEN COLLECTION AND SUBMISSION GUIDELINES:

Test Request Form: DCH-0558.

Specimen Collection Guidelines.

Transport Temperature: All temperatures acceptable.

SPECIMEN TYPE:

Dust: Individually packaged ASTM approved wipe.

  - Minimum Acceptable Volume: Area wiped not less than 0.25 sq. ft. and not greater than 2 sq. ft.

  - Container: Submit sample in polyethylene centrifuge tubes (50mL size).

Soil: Minimum sample size: 20 g avoiding excess twigs, stones, and other debris submitted in 50 mL polyethylene centrifuge tube.

Paint: Minimum sample 0.3 g paint in 50 mL polyethylene centrifuge tube.

Shipping Unit: Not Applicable.

SPECIMEN REJECTION CRITERIA:

Insufficient sample size.

Sample received in unsuitable container.

TEST PERFORMED:

Methodology: ICP-AES.

Turnaround Time: 1-7 days.

Testing is performed twice per week at the Bureau of Laboratories.

RESULT INTERPRETATION:

Adverse health effects result from environmental exposure to lead.

Dust Lead Clearance Levels:

  - 5 μg/ft2, floors.

  - 40 μg/ft2, interior windowsills.

  - 100 μg/ft2, interior window troughs.

  - 40 μg/ft2, exterior porch floors.

Soil Hazard Determination:

  - 400 μg/g bare soil in child play areas and gardens.

  - 1200 μg/g bare soils at the foundation/dripline and any other section of the yard with bare soil.

Paint Hazard Determination:

  - Paint levels equal to or exceeding 0.5% (5000 ppm) by weight.

FEES:

Not Applicable.

NOTES:

Results are reported for lead follow-up via StarLIMS to HL7 protocol for prompt action.

ALIASES:

Not Applicable.