HIV-1 Rapid Testing by Uni-Gold™ Recombigen® HIV-1input.doc

I.	Introduction & Principle

The standard laboratory HIV testing algorithm used in the United States consists of screening with an enzyme immunoassay (EIA) and confirmation of repeatedly reactive EIAs using a Western blot test or IFA. Results are typically reported within 48 hours to 2 weeks. Nearly one fourth of the estimated 900,000 HIV-infected persons in the United States do not know their HIV status.  This is because 30% of persons who tested HIV-positive during 2000 and 39% of persons who tested HIV-negative did not return for their test results. (HIV CT Client Record Report, 2000 U.S. Total; CDC unpublished data).

Uni-Gold™ Recombigen® HIV was designed as a rapid immunoassay and is intended to detect antibodies to HIV-1 in human whole blood (venipuncture and finger stick). Uni-Gold™ Recombigen® HIV uses proteins representing regions of the HIV virus.  If antibodies to HIV-1 are present in the sample, they combine with these proteins and a color reagent and this complex binds to the proteins in the test forming a visible pink/red band in the test region of the device adjacent to the word ‘Test’. The control line should always appear as a visible pink/red band in the control region of the device to indicate that the test device is functioning correctly.  A reactive result is indicated by a pink/red band in the test region of the device. A non-reactive result occurs in the absence of detectable levels of antibodies to HIV-1 in the specimen; consequently no visually detectable band develops in the test region of the device.

This test is to be used under the following CLIA complexity designation: For use with waived testing, for whole blood fingerstick and venipuncture samples only.

II.     Specimen Collection And Storage:

For venipuncture whole blood:  EDTA, acid citrate dextran (ACD) or heparin should be used as the anticoagulant. Other anticoagulants have not been tested and may give incorrect results.

Whole blood collected by fingerstick:

Whole blood samples collected by fingerstick should be used on the Uni-Gold™ Recombigen® HIV immediately after collection.

Whole blood collected by venipuncture:

Using standard phlebotomy procedures, collect a venipuncture whole blood specimen using a blood collection tube containing either EDTA, acid citrate dextran (ACD) or heparin.  Other anticoagulants have not been tested and may give incorrect results.

It is recommended that specimens be tested immediately but can be tested within 8 hours of collection if stored at ambient temperature (15°C- 27°C/ 59.0°F – 80.6°F). Grossly hemolysed or lipemic samples should not be used.

III.	Materials:

### MATERIALS PROVIDED

Each kit contains:

a)  20 Test Devices (individually pouched)

b)  Wash solution 5.0 ml

c)  20 Disposable Pipettes for use with venipuncture whole blood. To be used also with Controls (Catalog number 1206530)

d)  20 Disposable Fingerstick Sample Collection and Transfer Pipettes for use with fingerstick whole blood

e)  20 Subject Information Leaflets

f)   1 Package Insert

Materials required and available as an accessory to the kit

Uni-Gold™ Recombigen® HIV Kit Controls.  Catalog number 1206530.

Each pack of Kit Controls contains Positive Control, 1 vial (red cap), (0.5ml) and Negative Control, 1 vial (black cap) (0.5ml) and a package insert.

MATERIALS REQUIRED BUT NOT PROVIDED

Timer or stopwatch

Blood collection devices, for testing of venipuncture whole blood.

Biohazard disposal container

Disposable gloves

Disinfectant (i.e. 1:10 dilution of household bleach)

For Fingerstick samples the following additional materials are required.

●  Adhesive bandages

●  Lancet capable of producing a 50µl droplet

●  Sterile wipes and sterile gauze pads

WARNINGS

For in vitro diagnostic use

1.	Read the package insert completely before using the product.  The instructions must be followed carefully as not doing so may result in inaccurate results.

2. 	Before performing testing all operators must read and become familiar with the Standard Precautions for Prevention of Transmission of Human Immunodeficiency Virus, Hepatitis B Virus and other Blood-Borne Pathogens in Health-Care settings (1).

3. 	Use of the kit with specimens other than those specifically approved for use with this device may result in inaccurate test results.

- This test kit is CLIA-waived for use only with fingerstick whole blood and venipuncture whole blood samples.

5. 	Uni-Gold™ Recombigen® HIV is for diagnostic use only and is not to be used for screening donors of blood, plasma, cells or tissues.

- Perform test at room temperature (15°C – 27°C / 59.0°F – 80.6°F).

## Storage

Uni-Gold™ Recombigen® HIV device and Wash Solution must be stored between 2°C (35.6°F ) and 27°C (80.6°F).

Kit components are stable until expiration date when stored as directed.

If stored refrigerated, ensure that the pouched device is brought to room temperature [15°C (59.0°F) to 27°C (80.6°F] before opening.

Do not use beyond expiration date.

Do not freeze the kit.

Store the separately supplied Uni-Gold™ Recombigen® HIV Kit Controls at 2ºC-8ºC/35.6°F-46.4°F.

Handling Precautions

1.	Do not use any device if the pouches have been perforated.

2. 	Each device is for single use only.

3. 	Do not mix reagents from different kit lots.

4. 	Do not use the kit past the expiration date (this date is printed on the box).

5. 	Adequate lighting is required to read the test results.

6. 	Read results 10 minutes following the addition of Wash Solution. Do not read results more than 12 minutes following the addition of Wash Solution.

7. 	Lancets should be placed in a puncture resistant container prior to disposal.

IV.	    Safety:

  - Handle blood specimens and materials contacting blood specimens as if capable of transmitting infectious agents.

  - Do not drink, eat, or smoke in areas where specimens are being handled or testing is being performed.

  - Wear a lab coat, eye protection and disposable gloves while handling and testing blood specimens. Use a fresh pair of gloves for each client.

  - Dispose of gloves in a biohazard waste container after use.

  - Wash hands thoroughly after performing each test or utilize an alcohol-based hand sanitizer.

  - Dispose of all test specimens and materials, used in the test procedure, in a biohazard waste container.

  - Lancets should be placed in a puncture-resistant container prior to disposal.

  - The recommended method of disposal of biohazard waste is autoclaving for a minimum of 1 hour at 121ºC. Disposable materials may be incinerated. Liquid wastes may be mixed with appropriate chemical disinfectants.

  - Wipe all spills thoroughly with a solution of 10% bleach or other appropriate disinfectant. Fresh 10% bleach solution should be made daily.

  - NOTE: Do not autoclave solutions that contain bleach.  For additional information on biosafety, refer to "Universal Precautions for Prevention of Transmission of Human Immunodeficiency Virus, Hepatitis B Virus, and other Blood-borne Pathogens in Health-Care Settings".(2)

V.	Procedure:

TEST PROCEDURE AND INTERPRETATION FOR CLIA WAIVED SETTINGS

Test Procedure For Fingerstick Whole Blood

1.	Ensure that the Subject Information Leaflet has been given to the client.

2.	Allow the kit (unopened devices and Wash Solution) to reach room temperature [(15ºC (59.0°F) to 27°C (80.6°F)] for at least 20 minutes if previously stored in the refrigerator.  Once at room temperature remove the required number of Uni-Gold™ Recombigen® HIV devices from their pouches.

PERFORM ONLY ONE TEST AT A TIME.

3. 	Lay the device on a clean flat surface.

4. 	Label the device with the appropriate patient information.

5.   Sample collection and addition to the device (directions):

●Using an antiseptic wipe, clean the finger of the person being tested. Allow the finger to dry thoroughly or wipe dry with a sterile gauze pad.

● Using a sterile lancet capable of producing a 50µl blood let, puncture the skin just off the center of the finger pad.  Hold the finger downward. Apply gentle pressure beside the point of the puncture.  Avoid squeezing the finger to make it bleed. Wipe away the first drop of blood with a sterile gauze pad.  Allow a new drop of blood to form.  If blood flow is inadequate, the subject’s finger may have to be gently massaged at the finger base to produce a droplet of sufficient volume. Avoid ‘milking’ the finger.

●Collect the blood into the fingerstick sample transfer pipette provided   following the procedure presented below.

a. Hold the Pipette bulb gently in a horizontal position to the sample to be collected.  This is important, as the specimen may not be adequately drawn in the pipette if the pipette is held in a vertical position.

b. Place the tip of the Pipette into the sample, taking care not to squeeze the bulb.  Maintain this position until the flow of sample into the Pipette has stopped. The sample should fill to the mark on the Pipette. If sample is not collected to the mark, the Pipette should be safely discarded and another specimen should be collected from another finger by repeating the sample collection process.  The sample should be used immediately.

c. Squeeze the bulb until the sample is fully discharged into the Uni-Gold™ Recombigen® HIV sample port. Should the sample not fully discharge, cover the small opening at the mark on the Pipette with a gloved finger.   Then squeeze the bulb until the sample is fully discharged.  Allow the sample to absorb into the paper in the sample port. Ensure air bubbles are not introduced into the sample port.

d. Dispose the Pipette in biohazard waste.

6. 	After the blood sample has been completely absorbed, holding the dropper bottle of Wash Solution in a vertical position, add four (4) drops of Wash Solution to the Sample Port.

7.	Set the timer for 10 minutes and start timing the test.

8.	Read test results after 10 minutes but not more than 12 minutes incubation time.

9. 	Refer to the Test Results and Interpretation of Whole Blood Samples below.

10.Check for full red color in sample port. The sample port must contain red color for test to be valid.  A red/pink line must appear adjacent to the word control. A red/pink line may appear adjacent to the word test. If no red color is seen in the sample port repeat test with fresh device.

Test Procedure Venipuncture Whole Blood

1.	Ensure that the Subject Information Leaflet has been given to the client.

2. 	Allow the kit (unopened devices and wash solution) to reach room temperature [(15°C (59.0°F) to 27°C (80.6°F)] for at least 20 minutes if previously stored in the refrigerator. Once at room temperature remove the required number of Uni-Gold™ Recombigen® HIV devices from their pouches.  Perform no more than 10 tests at one time.

3. 	Lay the devices on a clean flat surface.

4. 	Label each device with the appropriate patient information / ID.

5. 	Draw up adequate sample to the first gradation on the pipette using one of the disposable pipettes included in the kit. Use only the pipette included in the kit and do not reuse.

6. 	Holding the pipette vertically over the sample port, carefully add one (1) free falling drop of sample. Do not add the full volume contained within the pipette. Allow the sample to absorb into the paper in the sample port. Ensure air bubbles are not introduced into the sample port. Discard the pipette in a biohazard waste container.

7. 	After the blood sample has been completely absorbed, holding the dropper bottle of Wash Solution in a vertical position, add four (4) drops of Wash Solution to the Sample Port.

8.	Set the timer for 10 minutes and start timing the test.

9. 	Read test results after 10 minutes but not more than 12 minutes incubation time.

10. Refer to the Test Results and Interpretation of Whole Blood Samples below.

11. Check for full red color in sample port.  The sample port must contain red color for test to be valid. A red/pink line must appear adjacent to the word control. A red/pink line may appear adjacent to the word test. If no red color is seen in the sample port repeat test with fresh device.

INTERPRETATION FOR WHOLE BLOOD SAMPLE

FOR A TEST TO BE VALID A CONTROL LINE MUST BE PRESENT AND THE SAMPLE PORT MUST CONTAIN FULL RED COLOR

Invalid Results

A.	Clinical Results Interpretation:

- Non-Reactive (Negative): These individuals are not infected unless they have had a recent (within 3 months) known or a possible exposure to HIV.  Recommend a retest for clients with a recent exposure pursuant to MDCH guidelines.

- Reactive (Preliminary Positive) Test: Further testing is always required to confirm a reactive screening test result.  See section VIII this procedure. The client must be instructed to avoid potential transmission of virus.  It is essential to explain:

      - The meaning of reactive screening test result in simple terms, avoiding technical jargon.

      - Emphasize the importance of confirmatory testing and schedule a return visit for confirmatory test results.

      - Underscore the importance of taking precautions to prevent transmitting infection to others while awaiting results of confirmatory testing.

- Refer to MDCH guidelines on counseling associated with rapid HIV testing for

- detailed information regarding requirements for disclosure of HIV test results.

VI.	Quality Control

A. Controls:

1.	Built-In Control Features:

The Uni-Gold™ Recombigen® HIV test has a built in procedural control that demonstrates assay validity.  A red / pink line adjacent to the word ‘control’ indicates that the test was performed correctly.  In addition, when using whole blood samples, there must be a red color in the sample port to validate the addition of the sample.  The control line will appear on all valid tests, whether or not the sample is Reactive or Non-Reactive (refer to the test results and interpretation sections).

External Quality Control:

Uni-Gold™ Recombigen® HIV Kit Controls (Product Code: 1206530) are available separately for use only with the Uni-Gold™ Recombigen® HIV test.  The Kit Controls are used to verify your ability to perform the test and interpret the test result.  The Positive Control will produce a Reactive test result and has been manufactured to produce a very faint Test line.  The Negative Control will produce a Non-Reactive test result (refer to the test results and interpretation section). Note that a red color at the sample port will not be seen if using the Uni-Gold™ Recombigen® HIV kit controls (Product Code: 1206530).

Run the Kit Controls under the following circumstances:

●	All new operators performing testing on client specimens

●	Each new kit lot

●	Whenever a new shipment of test kits is received

●	If the temperature of the test kit storage area falls outside of 2°C (35.6°F) to 27°C (80.6°F).

●	If the temperature of the testing area falls outside of 15°C (59.0°F) to 27°C (–80.6°F).

●	At periodic intervals as specified in the following section (B) this procedure, Frequency of Controls.

The Kit Controls must give the expected reactive or non-reactive results, otherwise the test results are not valid.  Refer to the Uni-Gold™ Recombigen® HIV Kit Control package insert for instructions on the use of these reagents. Contact Trinity Biotech Customer Service if the Kit Controls do not produce the expected results.

B.	Frequency of Controls:

- For sites testing more than 25 clients per day (high-volume), positive and negative controls are to be run once each day of testing. (A day is defined as a 24-hour period.  For example, if testing is conducted from 8 p.m. until 2 a.m. that would be considered one day.)

- When testing at an off-site or a mobile location, positive and negative controls are to be run once each day of testing, regardless of the number of clients being tested. (A day is defined as a 24-hour period.  For example, if testing is conducted from 8 p.m. until 2 a.m. that would be considered one day.)

- For all other sites, including sites testing fewer than 25 clients per day (low-volume), positive and negative controls  can be run once per week when testing is performed.

- In addition, positive and negative controls  must be run and  perform as expected before any patients are tested in the following situations:

  - With each new lot of test devices prior to placing them in service. (A test kit lot is defined as the number printed on the outside of the foil package containing the test devices).

  - Each new operator before testing any clients for the first time.

  - With each new shipment of test kits received (even if it is the same lot number in current use).

  - If there is any change in the conditions of testing (e.g. new location, lighting, temperature, etc.).

  - If the temperature of the test storage area has fallen outside of 2° to 27C (35 to 80F).

  - If the temperature of the testing area has fallen outside of 15°to 37°C (59° to 99°F).

  - Whenever two consecutive invalid test results are obtained on the same client.

C.	Quality Control Records

- Quality Control (QC) information is to be recorded on the appropriate QC Log Sheet. The information required includes name of test, site performed, date, time and person performing the test, lot numbers of all reagents, expiration dates of all reagents, expected results and observed results.

- Temperatures for kit storage area and the area where the test is performed are to be recorded each day.

D.	External Proficiency testing:

- Testing facilities must enroll in an external proficiency program. Acceptable programs include the Centers for Disease Control and Prevention MPEP, College of American Pathologists Proficiency or American Proficiency Institute.

- The laboratory director or technical consultant will review results. Acceptable performance is 80% or greater.

- If fewer than five samples were tested, acceptable performance is 100%.

- Unacceptable performance requires that the director or technical consultant approve a written corrective action plan, including but not limited to re-training.

- Sometimes a proficiency testing result is returned as “ungraded” by the          provider. This usually means that not enough laboratories obtained the result that was expected, and it may indicate a problem other than the proficiency of the site. However, the testing site must still compare its result

to the expected result, and the laboratory director must determine whether corrective action is appropriate.

- E.	Corrective Action:

- 1.	If the controls fail to yield the expected results, DO NOT perform any     patient testing until performance issues are resolved and expected results are obtained and recorded.

2.   Document the corrective action taken; see the Quality Assurance Manual

and Corrective Action Guidelines.

3.	All corrective action reports must be reviewed and signed by both the Site Coordinator and the Laboratory Director.

- VII.  Referral for confirmatory testing

- Whenever the Uni-Gold™ Recombigen® HIV-1 test result is reactive the testing site must have established procedures for referral of either test specimens or persons being tested for confirmatory testing.

- If specimens are collected on site, the site must establish procedures describing how to collect, label, process, store and document specimen transfer; transport the confirmatory test specimens to the site(s) where they will be tested; and obtain the confirmatory results to give the client/patients.

- Indicate on the specimen requisition form that the specimen is from an individual who had a reactive Uni-Gold™ Recombigen® HIV-1 rapid test result.  Affix the MDCH bright orange adhesive “OQ-R” label to both the specimen tube and the laboratory requisition.

- Sites not able to collect confirmatory test specimens must have a procedure in place for referring persons to another site to obtain this testing.

- VIII.  Confirmatory Testing Protocols

- All Uni-Gold™ Recombigen® HIV-1  reactive (preliminary positive) results must be followed up with a Western blot (WB) for confirmation.

- Confirmatory testing can be done on blood (plasma, serum or dried blood spots) or oral fluid specimens.

- The reference laboratory will perform enzyme immunoassay (EIA) screening tests prior to the Western Blot. However, regardless of the EIA result, a reactive Uni-Gold™ Recombigen® HIV-1 must always be followed by Western Blot testing.

- IX.	Follow up testing when confirmatory testing is negative (discordant result)

- A reactive (preliminary positive) Uni-Gold™ Recombigen® HIV-1 test and a negative or indeterminate Western Blot confirmatory test is considered a discordant result. Always follow the MDCH guidelines and procedures for recommending retesting and reporting of discordant results.

- The client should be re-tested as soon as possible to rule out pre-analytical  (e.g. mislabeling), analytical (e.g., problem with kit), or post-analytical (e.g., wrong result reported on client) errors. The client should have a new blood specimen collected, and screening and confirmatory testing should be repeated.

- If the results are still discordant and the possibility of error and/or technical problems has been ruled out, collect a new blood specimen in four weeks and repeat screening and confirmatory tests.

- X.	Limitation of Method:

1.	Uni-Gold™ Recombigen® HIV must be used in accordance with the instructions   in this procedure to obtain an accurate result.

2. 	Uni-Gold™ Recombigen® HIV is designed to detect antibodies to HIV-1 in undiluted whole blood (venipuncture and fingerstick). For venipuncture whole blood: EDTA, acid citrate dextran (ACD) or heparin should be used as the anticoagulant. Other anticoagulants have not been tested and may give incorrect results. Other body fluids may not give accurate results and must not be used.

3. 	Immunosuppressed or immunocompromised individuals infected with HIV-1 may not produce antibodies to the virus.  Testing with any kit designed to detect antibodies may give negative results in this incidence and would not be a reliable test method for such patients.

4. 	The intensity of a line at the "Test" region is not an indication of the level of antibody in the specimen.

5. 	A Reactive result by Uni-GoldTM Recombigen® HIV suggests the presence of anti-HIV-1 antibodies in the specimen. Uni-GoldTM Recombigen® HIV is intended as an aid in the diagnosis of infection with HIV-1. AIDS and AIDS-related conditions are clinical symptoms and their diagnosis can only be established clinically.

6. 	Reading test results earlier than 10 minutes or later than 12 minutes may give incorrect results.

7. 	A Non-Reactive result with Uni-GoldTM Recombigen® HIV does not exclude the possibility of infection with HIV. A false negative result may occur in the following circumstances:

●	Recent infection.  Antibody response to a recent exposure may take several months to reach detectable levels.

●   The test procedure has not been correctly followed.

●	Antibodies to a variant strain of HIV-1 in the patient that does not react with specific antigens utilized in the assay configuration.

●    Improper specimen handling.

●	Failure to add sample.

8.	A person who has antibodies to HIV-1 is presumed to be infected with the virus,    except when a person who has participated in an HIV vaccine study may develop antibodies to the vaccine and may or may not be infected with HIV. Clinical correlation is indicated with appropriate counseling, medical evaluation and possibly additional testing to decide whether a diagnosis of HIV infection is accurate.

- XI. References:

- 1.  CDC. Universal precautions for prevention of transmission of human       immunodeficiency virus, hepatitis B virus and other blood-borne pathogens in

- health-care settings. MMWR 1988; 37(24):377-388.

- 2.  Quality Assurance Guidelines for Testing Using Rapid HIV-1 Antibody Test. U.S. Department of HHS and CDC. April 2005.

- 3.   Package Insert #045-146 for Uni-Gold™ Recombigen® HIV, June 2005. Trinity        Biotech USA, St. Louis MO 63114.

This material reviewed and approved for use without modification:

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

RL.11.02

Rev. 1/9/2007