Urine Pregnancy Test - RefuAH hCG

- Purpose:

The RefuAH hCG Assay is a chromatographic immunoassay (CIA) for the qualitative detection of human chorionic gonadotropin (hCG) in urine specimens to aid in the early detection of pregnancy.  Human chorionic gonadotropin (hCG) is a hormone produced by the developing placenta during pregnancy.    The test utilizes a combination of monoclonal and polyclonal antibodies to selectively detect elevated levels of hCG in urine soon after conception.

The assay is conducted by adding urine to the specimen well of the test device and observing the formation of colored lines.  The urine sample migrates via capillary action along the membrane to react with the colored conjugate.  Positive specimens react with the specific colored antibody conjugates and form a colored line at the test line region of the test device.  Absence of this colored line suggests a negative result.  To serve as a procedural control, a colored line will always appear at the control line region if the test has been performed properly.  The test is completed within 3 minutes.

- Specimen:

A.	Urine specimens should be collected in a clean, dry container such as a urine collection cup.

B.	Specimens may be collected at any time of the day.  First morning urine samples will normally contain the highest levels of hCG.

C.	If not tested immediately, urine may be stored at refrigerated at 36─46oF (2─8oC) for up to 48 hours.  Samples must be brought to room temperature before testing.

D.	For prolonged storage, specimens may be frozen and stored below -20oC.  Frozen specimens should be thawed and mixed before testing.

E.	Urine samples exhibiting visible precipitates should be filtered, centrifuges, or allowed to settle, to obtain aliquots for testing.

F.	Specimen rejection criteria:  Samples of unknown age or in unapproved containers should be rejected, and a fresh sample collected.

SAFETY NOTICE: Human specimens may harbor infectious agents.  Use standard (universal) precautions when working with these materials including gloves and eye protection.

- Materials

- A.	RefuAH hCG Prenancy Test (Sacks Medical Corporation, cat # W1SMC-C).  25 test strips per kit,

- 1.	Test devices

- 2.	Disposable specimen droppers

- 3.	Package insert

- B.	Required materials not supplied with kit

    - Specimen collection containers

    - Timer or watch that measures minutes and seconds

- External controls

- 1.	A positive hCG control (containing 25-250 mIU/mL hCG) and a negative control containing “0” mIU/mL hCG) must be used.  Several vendors are available which supply these controls.

- a.	Quidel hCG controls (Quidel Corporation, cat. # 00272)

- b.	Chek-Stix controls (Bayer, cat. # 1364)

- IV.	Storage and Stability

- The test kit is to be stored either refrigerated or at room temperature (40-86oF or 4-30o C) out of direct sunlight.  The test device is stable through the expiration date printed on the sealed pouch.  The test device must remain in the sealed pouch until use.  DO NOT FREEZE.  Do not use beyond the expiration date.

V.	Quality Control

- External Controls

1.	External positive and negative controls are to be performed on each new lot of test kits or with each new shipments of test kits.

2.	If the controls do not yield the expected results, the kit must be removed from service and the action noted in the corrective action section of the QC log sheet.

3.	The responsibility for performance of external controls must be rotated among all staff performing testing.  The designation of a specific individual to perform all QC activities in the clinic is not permitted.

B.	Built-in Procedural Control

1.	A procedural control is included in the test.

2.	A pink line appearing on the control region (labeled “C”) of the test device indicates that proper volume of fluid was absorbed into the test strip and capillary flow occurred.

3.	If the Control line does not develop within 3-5 minutes, the test result is invalid.

4.	Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure.  Review the procedure and repeat the test with a new test device.  If the problem persists, discontinue the test kit immediately and contact Sacks Medical Corporation at 1-800-521-1635.  The site coordinator and laboratory director or technical consultant must also be notified.

5.	The result of the procedural control must be documented for the external positive control and the external negative control.  Corrective action must be initiated and documented whenever the blue Control line does not develop as expected.

C.	Prior to using a new shipment or lot number of RefuAH hCG, the Positive Control and Negative Control should be tested and shown to yield the expected  results.  Upon observing the expected results, the kit is ready for use with patient specimens.

A.	The Quality Control Log Sheet should include:

1.	Device name and manufacturer

2.	Date package, or kit, opened

3.	Lot number and expiration date of pregnancy testing device

4.	Lot number and expiration date of each control reagent

5.	Results of:

a.	Positive Control

b.	Negative Control

c.	Procedural Control

6.	Initials of staff person conducting quality control tests.  The site supervisor and laboratory director must review and sign all QC forms on a quarterly basis.

B.	Record the last number and expiration date of the pregnancy test device on the daily clinical worksheet.

C.	Store records for two years.

- Method

  - Review specimen collection instructions.

  - Allow the urine specimen and controls to equilibrate to room temperature prior to testing.

  - Remove the “reaction pack” from its foil wrapper by tearing along the “splice”.

  - Place the test device on a clean and level surface.

  - Fill the urine dropper with urine and hold the dropper vertically and transfer four(4) full drops or urine(approximately 100 μl) to the specimen well (S) of the test device, and then start the timer.  Avoid trapping air bubbles in the specimen well (S).

  - Wait for the pink line(s) to appear.  Positive results may be read as early as three (3) minutes.  It is important that the background is clear before the results are read.  Do not interpret results after five (5) minutes.

- NOTE:  Early stages of pregnancy may show positive results nearer to five (5) minutes due to low concentrations of hCG.

- VI.	Results

- The RapidVue hCG pregnancy test detects urinary hCG at concentrations of 25 mU/ml or greater.  This is indicated by the appearance of a pink line in the test region.

- A.	Negative: At 3-5 minutes, only one pink line appears in the control region (C).  No apparent pink line appears in the test region (T).

- Positive: At 3-5 minutes, two distinct pink lines appear.  One line should be in the control region (C) and another line should be in the test region (T).

- C.	Invalid: At 3-5 minutes, if no bands appear, or a test band appears without a control band, the result is invalid and thest should be repeated using a new device. All failed tests must be documented and the appropriate corrective action must be initiated and documented.

- NOTE:  Do not interpret results after 5 minutes.

- Procedural Notes

- A.	Put a new lot of test strips into use if the control line fails to appear on more than two tests from the same lot number.  Deterioration of the test strip may have occurred.

- B.	It is important to interpret negative test results at the end of the reaction time (3 minutes).  The manufacturer package insert states that results are not to be interpreted after 5 minutes.

- D.	Results may be considered positive as soon as the test band turns color as long as it is within the 3-5 minute interval.

  - The shade of color which develops on the test band may vary from pink to red depending upon the amount of hCG present in the sample.  However neither the intensity of color nor time of color change can be interpreted as an accurate quantitative measurement of hCG.

  - Weakly reactive positive or negative test results in patients suspected to be pregnant should be confirmed by retesting a fresh early morning urine specimen obtained 42 to 72 hours later or performing a quantitative hCG assay.

- Limitations of the procedure

  - Proteinuria, hematuria, gross bacterial contamination, and detergents may cause an increase in HCF in urine.

  - False Positive:   patients with trophoblastic diseases, hydatifiform mole, or choriocarcinoma may yield false positive results.  Patients with ovarian and testicular teratomas have also been reported to excrete large quantities of hCG.

  - False Negative:    diluted or low specific gravity urine and conditions that may cause denaturation of hCG (e.g., pH, temperature, contamination with heavy metals, and so forth) may yield false negative results.

  - Samples with hCG concentrations less than 25 mIU/mL will be detected as negative.

- Expected Values:

  - Healthy men and healthy non-pregnant women do not have detectable hCG in urine by the RapidVue hCG test.

  - In approximately seven (7) days after conception, hCG concentrations of 5 – 50 mIU/mL in serum and urine will appear.  This level of hCG will reach 100 mIU/mL within 60 to 80 days and will be followed by a gradual fall to a mean level of about 20,000 mIU/mL after 120 days.

- References:

  - RefuAH Package Insert (Sacks Medical Corporation)

Technical support: 1-800-521-1635.

***********************************************************************************

This material reviewed and approved for use without modification:

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

Review Date/Signature: ______________________________________________________________

RL42.02

Rev.  8/2011